## Slide 1: Single Crystals Grown for radiation detection
Diameter controlled Ge single crystal Size: 50 mm f and 150 mm L
7/7/2022
LiI:Eu Crystals 50 mm f and 50 mm L
LaBr3:Ce Crystals 50 mm f and 50 mm L
YAlO3:Ce Crystal (size 30 mm f and 40 mm L)
Gd3Ga3Al2O12:Ce Crystal (size 30 mm f and 40 mm L)
CsI:Tl Crystal (size 75 mm f and 100 mm L)

## Slide 2: YAP crystal used to fabricate E-T detector for SEM

E-T Detector (for SEM) fabricated using the grown YAP crystal
Images recorded using SEM detector

## Slide 3: Development of 1D (16x1 matrix) and 2D (8x8 matrix) array of Pixelated CsI:Tl crystals for imaging applications
1 D array of Pixelated CsI:Tl crystals (16x1) dimension 3 mmx1.4mm with separation of 100 micron and mounting on silicon diode photodetectors for the X-ray Baggage Scanner project (TPD, ED, ECIL)
2D Pixelated CsI:Tl Single crystal mounted on PSPMT.
Spatially resolved image of a gamma source kept in a phantom.
The Small Organ Imaging Gamma camera is for detection of Sentinel node of Breast Cancer by uptake of 131I tracer isotope.(TPD & SSSIHL)
The response of each pixel with X ray is found to be comparable with Saint-Gobain made detectors

## Slide 4
Development of Poshwisch detectors for discrimination of different radiations
Phoswich Geometry for neutron measurement from spontaneous fission of Pu
The pulse height spectra measured in GGAG crystal only for samples having Pu
2D Pulse Shape Discrimination plot of phoswich
A novel versatile phoswich detector consisting of two single crystals to discriminate various kinds of radiations: US patent No. 16/358427

## Slide 5
High energy Gamma Spectrometer based on BGO crystal
Phoshwich Detector

## Slide 6
7/7/2022
Gain: 15
Shaping: 3µs
Bias: +1000 V
Active Volume: 3 cc
HPGe diode developed at TPD from High Purity Ge Crystal from Umicore
137Cs 
HPGe Detector Development : TPD (Sensor), EmA&ID (Cryostat) & ED (Electronics)
2022

## Slide 7: Comparison with 8cc HPGe Commercial detector
Peak Position (KeV) | CTL P-Type, Planar, 3 cc Bias: +1250V FWHM (KeV) | FCD P-Type, Planar, 8 cc Bias: -1800V FWHM (KeV)
121.8 | 0.80 | 0.55
244.8 | 0.93 | 0.72
344.3 | 1.00 | 0.84